Status & Plans for

MPA-ALMY Sensor R&D at MPI Munich

Volker Bartheld, Sofia Chouridou,

Hubert Kroha, Stefan Schael

# Max-Planck-Institut für Physik

# Föhringer Ring 6

D-80805 München

- Introduction

- Prototype production & test

- Optical properties of the MPA-ALMY sensors

- Sensor stability & ageing

- Summary & conclusions

Myonchamber alignment in the ATLAS endcap:

Requirements for the alignment system:

- Large active area of the sensors, 20mm  20mm for MPA-ALMY sensors

- Resolution in the order of 5 m

- Sensor transparency in the order of 90%

- Minimal laser beam distortion by the alignment sensor

- Long term sensor stability, (10 years)

- Linear and stable laser beam over 20 m distance

MPA-ALMY Sensors

- transparent a-Si:H sensors

- size: 20 20 mm

- 64 strips in X- and Y- direction

- 150 such sensors have been build and tested

Testsetup:

## Laser: 780 nm wavelength, power 3-5 mWatt

Positionresolution: ca. 5 m in X- and Y-direction

transmission: ca. 81%

##### System of 6 sensors: 1. Sensor signal height 2000 ADC counts

6. Sensor signal height   500 ADC counts

Sensitivity varies by a factor of 2 between the sensors

Test results for sensor vs150:

Summary for the prototype production

Requirements for “good” sensors:

- Position resolution over the full surface: (pos.) < 5 m

- Gaussian resolution distribution

- Deflection angle variation (def.) < 5 rad

- Gaussian deflection angle distribution

From the 117 sensors produced & tested:

- 117 have (pos.) < 5 m

- 56 have a Gaussian resolution distribution

- 1 (vs150) has (def.) < 5 rad

- 1 (vs150) has a Gaussian deflection angle distribution

Conclusion:

All sensors that have been tested so far would need additional corrections for the deflection angle variation to be used in a multipoint alignment system. 50% of them would also need corrections for the position measurement.

What is the reason for the observed problems ?

- thickness variations of the a.-Si:H  ?
transmission: T=(825)%   T/T  =   6%
absorption: A=(185)%  A/A = 28%

if we assume: Intensity I = c A we expect:
I/I = 28% but we measure I/I = 6%

- interference pattern ?
i.e. the different surfaces are not parallel

Testsetup:

From the incomming light:

- ca. 15% are reflected

- ca. 5% are absorbed

- ca. 80% transmitted

- Reflection & Transmission are 100% anti-correlated, as expected

- thickness variations in the a.-Si:H could not explain these data

- R+T+A = 1,  R+T=const  variations in A are small

An Anti-reflex-coating of the glas substrat can solve the problem

thickness:  d = /4 = 200 nm

material : n1=n0ns,    for example MgF2  with n = 1.38

with coating					              without coating

with coating					                 without coating

- the thickness of the ITO, a.-Si:H and glass have to be optimized for the optical properties of the system.

- therefore precise measurements of the absorption- and reflection-coefficient and of the thickness of the various materials are needed.

- this will lead to a new sensor design including an anti-reflection coating on the glass substrate.

- we are waiting for the material probes to be delivered by Heimann Optoelectronics. JENA-Optic will do the measurements and optimize the design.

Long term stability of the sensors

- within 1 hour the sensor loses 15% of its sensitivity

- afterwards the sensor is stable

This effect is known since 1977 as the Staebler-Wronski Effect from the study of solar cells. Up to know it is not fully understood.

Simple models describing the creation of defects in the a.-Si:H are in good agreement with the experimental data.

We use I = 1 mWatt/mm2  100 mWatt/cm2

Solutions:

- Reduce the laser power, i.e. build a new readout electronic

- Pre-aging of the total sensor

- Don’t use a continuos laser but synchronize the laser pulses with the readout electronics (ms).

Monte Carlo simulation:

Laser stability

Position variation reduced from 20 m/1C to 1 m/1C !

CCD camera					              MPA-ALMY Sensor

#### Summary and Conclusions

- 150 MPA-ALMY Sensors have been produced and tested.

- The position resolution is ca. 5 m over the full sensor surface.

- The optical properties of the sensors need to be optimized in order to use them in an MULTIPOINT-Alignment System.

- The long term stability of the sensors need to be demonstrated.

- A lot of effort is needed to get a long term linear and stable laser beam.

- We have ideas how to solve all technical problems observed so far, but a lot of R&D work is needed before a reliable system can be presented.