Status & Plans for

MPA-ALMY Sensor R&D at MPI Munich

Volker Bartheld, Sofia Chouridou,

Hubert Kroha, Stefan Schael

# Max-Planck-Institut für Physik

# Föhringer Ring 6

D-80805 München

- Introduction

- Prototype production & test

- Optical properties of the MPA-ALMY sensors

- Sensor stability & ageing

- Summary & conclusions

Muonchamber alignment in the ATLAS Endcap:

Requirements for the alignment system:

- Large active area of the sensors, 20mm  20mm for MPA-ALMY sensors

- Resolution in the order of 5 m

- Sensor transparency in the order of 90%

- Minimal laser beam distortion by the alignment sensor

- Long term sensor stability, (10 years)

- Linear and stable laser beam over 20 m distance

MPA-ALMY Sensors

- transparent a-Si:H sensors

- size: 20 20 mm

- 64 strips in X- and Y- direction

- 150 such sensors have been build and tested

Testsetup:

## Laser: 780 nm wavelength, power 3-5 mWatt

Position resolution: ca. 5 m in X- and Y-direction

transmission: ca. 81%

##### System of 6 sensors: 1. Sensor signal height 2000 ADC counts

6. Sensor signal height   500 ADC counts

Sensitivity varies by a factor of 2 between the sensors

Test results for sensor vs150:

Summary for the prototype production

Requirements for “good” sensors:

- Position resolution over the full surface: (pos.) < 5 m

- Gaussian resolution distribution

- Deflection angle variation (def.) < 5 rad

- Gaussian deflection angle distribution

From the 117 sensors produced & tested:

- 117 have (pos.) < 5 m

- 56 have a Gaussian resolution distribution

- 1 (vs150) has (def.) < 5 rad

- 1 (vs150) has a Gaussian deflection angle distribution

Conclusion:

All sensors that have been tested so far would need additional corrections for the deflection angle variation to be used in a multipoint alignment system. 50% of them would also need corrections for the position measurement.

What is the reason for the observed problems ?

- thickness variations of the a.-Si:H  ?
transmission: T=(825)%   T/T  =   6%
absorption: A=(185)%  A/A = 28%

if we assume: Intensity I = c A we expect:
I/I = 28% but we measure I/I = 6%

- interference pattern ?
i.e. the different surfaces are not parallel

Testsetup:

From the incomming light:

- ca. 15% are reflected

- ca. 5% are absorbed

- ca. 80% transmitted

- Reflection & Transmission are 100% anti-correlated, as expected

- thickness variations in the a.-Si:H could not explain these data

- R+T+A = 1,  R+T=const  variations in A are small

Thickness variations in the a.-Si.:H would lead to variations in the deflection angle, i.e. the laser beam distortions would scale with the distance from the sensor. This can be tested experimentally.

The observed laser beam distortion is independent of the distance between the CDD-camera and the MPA-ALMY Sensor. Therefore the observed laser beam distortion can not have its origin in thickness variations of the a.-Si.H.

An Anti-reflex-coating of the glass substrate can solve the problem

thickness:  d = /4 = 200 nm

material : n1=n0ns,    for example MgF2  with n = 1.38

with coating					              without coating

with coating					                 without coating

- the thickness of the ITO, a.-Si:H and glass have to be optimized for the optical properties of the system.

- therefore precise measurements of the absorption- and reflection-coefficient and of the thickness of the various materials are needed.

- this will lead to a new sensor design including an anti-reflection coating on the glass substrate.

- we are waiting for the material probes to be delivered by Heimann Optoelectronics. JENA-Optic will do the measurements and optimize the design.

Long term stability of the sensors

- within 1 hour the sensor loses 15% of its sensitivity

- afterwards the sensor is stable

This effect is known since 1977 as the Staebler-Wronski Effect from the study of solar cells. Up to know it is not fully understood.

Simple models describing the creation of defects in the a.-Si:H are in good agreement with the experimental data.

We use I = 1 mWatt/mm2  100 mWatt/cm2

Solutions:

- Reduce the laser power, i.e. build a new readout electronic

- Pre-aging of the total sensor

- Don’t use a continuos laser but synchronize the laser pulses with the readout electronics (ms).

Monte Carlo simulation:

#### Summary and Conclusions

- 150 MPA-ALMY Sensors have been produced and tested.

- The position resolution is ca. 5 m over the full sensor surface.

- The optical properties of the sensors need to be optimized in order to use them in an MULTIPOINT-Alignment System.

- The long term stability of the sensors need to be demonstrated.

- A lot of effort is needed to get a long term linear and stable laser beam.

- We have ideas how to solve all technical problems observed so far, but a lot of R&D work is needed before a reliable system can be presented.